Technical Specification of Fluorescent Microscope

A. Basic equipment characteristics

1. Trinocular upright microscope with camera attachment with bright field, dark field (upto 40X), phase contrast & fluorescence.

B. Facilities/functionalities available

1. Microscope Body: Antifungal treated trinocular upright microscope with camera attachment having infinity corrected optics with magnification range of 40X - 1000X for observation. Should have trinocular eyepiece tube having light distribution for simultaneous observation (100:0, 20:80, 0: 100) with eyepiece magnification of 10X / F.O.V of 25mm or more having dioptre adjustable (+/-5) on both eyes & should be anti-fungus type. Body with status display to view the optical configuration.

2. Illumination: The microscope must be provided with a high-performance LED illumination with power consumption 3W-5W or higher, life time of minimum 50,000 hrs. Should have uniform light distribution throughout the observation field and feature to adjust the light intensity as per change in magnification. Should have the latest technology featuring energy saving mode, ECO mode, sleep mode, light intensity management (LIM) function.

3. Condenser: Universal turret condenser for bright field, dark field and phase contrast.

4. Objective: High performance objective suitable for bright field, dark field (upto 40X), phase contrast & fluorescence.

a. Plan Fluor or similar 4X N.A 0.13

b. Plan Fluor or similar 10X N.A 0.30, Fluorescence, Phase contrast

c. Plan Fluor or similar 20X N.A 0.50, Fluorescence, Phase contrast

d. Plan Apochromat 40X N.A 0.95, Fluorescence

e. Plan Fluor 60X N.A 0.7, Fluorescence, Phase contrast with correction collar ring

f. Plan Apochromat 100X oil N.A 1.45, Fluorescence

g. Spring loaded object marker (Qty. 02) with each unit

h. Polarizing accessories should be provided

i. Higher NA will be preferable

5. Objective turret: Should have six positions revolving objective turret inward facing with intelligent / coded function having analyser slot. It should detect the objective magnification and adjust the scale bar while imaging.

6. Stage: Rust free rectangular mechanical stage having a capacity to hold two slides at a time. It should have features to reduce damage of the slide and mechanical depreciation of the stage. Upper limit stopper should be there to protect the slide.

7. Fluorescence illumination: System should have LED light sources (Wavelength 385nm, 475nm, 550nm, 621nm) (same make), software / PC control having long life (>20,000 hrs.). Should have 06 position fluorescence filter rotating turret with wavelengths and filters to support DAPI, FITC & TRITC.

a. Three sets of Fluorescence filter cube wavelength should be standard band-pass filter for DAPI, FITC and TRITC.

b. One set of Fluorescence filter cube wavelength should be long-pass filter:

FITC:- EX : 510-560 DM : 575 BA : 590

TRITC:- EX : 450-490 DM : 505 BA : 520

DAPI:- EX : 362-396 DM : 415 BA : 432-482

8. Digital Camera system: Scientific microscopic digital colour camera system with high resolution CMOS camera of 20 megapixels or more having:

a. size of 34mm X 22mm or better

b. more than 1 inch

c. live display 60-65 frame per second at full HD Resolution (1920 x 1080)

d. 10-12 bit depth

e. pixel size 5.5 μm x 5.5 μm or more

f. exposure time 100 μsec - 120 sec

g. photometry mode, live cell imaging, binning feature

h. high speed USB 3.2 PC interface

i. should be capable to capture fluorescence images/signal

j. required adapter, relay lens should provide to get image as microscope observation

k. should work on colour and monochrome image (bright field, phase contrast, dark field, polarizing, fluorescence, etc.)

9. Teaching Head Attachment (FOR ONE UNIT): Penta-head attachment along with above microscope specs: (1+4 heads) with binocular eyepiece tube, eyepiece of 10X magnification, LED arrow pointer, 360-degree rotatable pointer, dual-colour LED adjustment function, LED intensity light adjustment function.

10. Table: Hard granite top table with sufficient space for Penta-head microscope along with five height adjustable revolving stool with back rest having good quality cushion.

11. Microscope, optics, LED fluorescence, camera, image analysis software and accessories should be of same manufacturer for compatibility & future upgradability.

12. A compatible 2KVA UPS for the microscope and its accessories with at least 30-40 min back up time should be included.

13. A set of microscope stage micrometer and eyepiece reticle should be included for each microscope.

C. Information management

1. Computer: (Branded like HP, Dell, Lenovo, HCL) with latest generation i7 processor, 2TB or more SSD, 32GB or higher RAM, 2GB or more graphics card, wireless keyboard & mouse, at least 26” LED screen, latest Windows operating system (Windows 11 professional or higher), original Microsoft office with lifetime license and compatible 1KVA UPS.

2. Software: Licensed imaging software (same make) for 4-D imaging, image analysis, live cell imaging, AVI, multichannel & multipoint image capturing, intensity measurements, auto counting, auto measurement, volume view, 2D - 3D large image stitching (after stitch smooth image should be formed), line profile, Z stack, spatial measurements such as length/ width/area/perimeter, flu. channel mixing / un-mixing, pixel classifier, ROI, multi-user environment, report generation, etc. Software should be further upgradable to advance features such as AI functions, etc.

D. Certification & demonstration

1. The equipment must possess quality certifications such as FDA, or CE, or BIS to ensure compliance with regulatory standards.

2. Only principal companies or authorized distributors from principal companies should quote. Quotations from non- authorized distributors, re-sellers/ sub-authorization will not be considered & entertained. Features in the quotations should be sustained with proper principal company catalogue.

3. Should operate under a Quality Management System which complies with the requirements of ISO 9001: 2015 for design, manufacture and services.

4. Technical Compliance Statement in tabular form (mentioning bid specifications, product specifications, and compliance status) must be prepared and uploaded in the technical bid.

5. A demonstration of the equipment's performance in the presence of an authorized committee is mandatory in AIIMS-CAPFIMS campus following technical qualification.

6. User reserves the right to evaluate the software function, its UI.

7. Should any changes or disputes arise between the principal firm and the distributor, the department must be notified, and the principal firm is responsible for ensuring continued service provision in accordance with the established terms and conditions.

8. The technical bid must include at least five installation references of the same/older model with certificates of satisfactory performance for a minimum of five years. These references should be from institutes of national importance (INIs). Each reference must be signed by the Head of Department (HOD) or lab in-charge, along with their contact details.

9. Original literature for the equipment and consumables must be provided in English to facilitate understanding and compliance.

10. Specification must be supported by valid data sheet & information must be available on the website.

11. Documentation for Operation Qualification (OQ), Performance Qualification (PQ), and Installation Qualification (IQ) should be submitted at the time of tender. Validation of various parameters must be completed to the user's satisfaction during the installation process.

E. Electrical & other hardware requirements

1. Power Requirements: The equipment must operate on a power supply of 220 to 240 V AC at a frequency of 50 to 60 Hz.

2. The firm is required to supply a stabilizer cum UPS with minimum one hour back up, along with the equipment at no additional cost.

F. Consumables

1. Any consumable or accessory essential for the performance of the equipment that is not explicitly mentioned in the rate contract must be supplied by the firm at no additional cost.

2. The final bid evaluation will take into account the cumulative cost, which includes the equipment cost, and Comprehensive Maintenance Contract (CMC) costs for the 3rd to 10th years.

3. All rates must be quoted in Indian currency to ensure clarity and consistency in financial evaluations.

4. The bidder must quote the rates of consumables separately which shall be frozen for 10 years and considered for bid ranking. These consumables may be ordered as and when required by the end user department.

G. Troubleshooting and breakdown

1. In-house Service Engineers from principal company/ or authorized agents should be available in Delhi-NCR India on one-day notice basis in case of emergency (this will apply for that equipment wherein this clause is not built into the specifications). If the downtime is more than 24 hrs, penalty shall be levied as per institute rules.

2. Complaint points should be given (in the sheet) in order of the specifications’ serial order. Complaint points should be highlighted in the company catalogue with page number.

3. Must include: User’s list preferably containing initials with telephone numbers and email address.

4. Distributor should submit an undertaking from the manufacturer regarding the responsibility of maintenance in case of change of distributor/ Indian agent or merger or acquisition.

H. Warranty & CMC

1. The equipment must be brand new and not refurbished.

2. The supplier is required to provide original documentation that verifies the date and location of the equipment’s manufacture.

3. The comprehensive warranty must cover the replacement of all parts, including those subjects to wear and tear, regardless of whether they affect the routine functioning of the equipment, for a period defined by institute policy. The CMC (Comprehensive Maintenance Contract) rates for the subsequent period defined by institute policy, including accessories, UPS, and computer systems, must be quoted at the time of tender. If the company is unable to repair the equipment for any reason, a replacement with similar specifications and throughput must be installed free of cost, under the same terms and conditions, until the 10-year period is fulfilled.

4. A complete list of accessories and spare parts must be provided, including a comprehensive set of spare parts. The make, model, rating, description, price, and quantity of each item should be furnished separately. The rate contract for these spare parts will be valid for a period of 10 years.

5. All upgraded and new softwares should be provided free-of-cost for lifetime (warranty & CMC)

6. Warranty will be for 2 years and CMC will be for a period of 8 years following the warranty period